Development of control strategies for earthquake reproduction on shaking table

Afonso Henrique
Earthquake Engineering and Structural Dynamics Unit (NESDE/DE)
Laboratório Nacional de Engenharia Civil (LNEC)          
Lisbon, Portugal
ahenrique@lnec.pt
Miguel Ayala Botto
IDMEC, Instituto Superior Técnico, Universidade de Lisboa, Portugal

ayalabotto@tecnico.ulisboa.pt
Fernando Oliveira
Modelling and Rock Mechanics Unit (NMMR/DBB)
Laboratório Nacional de Engenharia Civil (LNEC)          
Lisbon, Portugal
fvoliveira@lnec.pt

Abstract—On this communication a new methodology is proposed for controlling shaking table systems in reproducing earthquakes. The system is identified online (closed‑loop) using the two‑stage method, and the controller is synthesized through an optimization algorithm that minimizes the squared error between the target and output response spectra. The validation of this methodology is achieved on a small-scale shaking table whose PID-F and LQI optimal controllers are used to evaluate the system performance. Numerical simulations using the identified physical model subjected to eight earthquakes show that both controllers outperform the OLI strategy currently used at LNEC, with the LQI controller achieving the best system performance. These findings partially validate the effectiveness of the proposed methodology, whose physical implementation will be carried out to confirm its performance on the experimental system.

Keywords— Earthquake Engineering, Control Systems, Closed-loop Identification,

# Introduction

Earthquakes are among the most catastrophic natural events, frequently causing severe loss of life, extensive property damage, and significant socio‑economic disruption. One of the most effective ways to reduce seismic risk is by decreasing the structural vulnerability of the built environment, which can be assessed through experimental techniques such as shaking‑table testing.

Shaking‑table tests consist of mounting a physical model on a stiff platform that is driven so as to reproduce the desired ground‑motion time history (reference signal). This experimental technique was introduced in the 1940s and became widespread during the 1960s. Early systems relied on mechanical profiled cams to generate the excitation, but servo‑hydraulic systems soon became the standard solution ([1],[2],[3]). At the Laboratório Nacional de Engenharia Civil (LNEC) a Goodmans model VG-108 electrodynamic shaker attached to a movable table was installed in the 60s to perform seismic tests [4]. a uniaxial shaking table (ST1D) was developed and installed, consisting of a 3 m × 2 m horizontal steel platform and a vertical platform body driven by a hydraulic actuator supplied by a dedicated oil‑hydraulic circuit ([5], [6]). During this period, adaptive control strategies were developed, implemented, and validated both on a small scale shaking table and on the ST1D, for tracking reference earthquake spectrum (frequency domain), iteratively ([8],[9]). Increasingly demanding requirements for reproducing seismic excitation along all three directions motivated the development of the triaxial shaking table (ST3D), which has been installed and operational at LNEC since the mid‑1990s [4]. The reproduction of earthquakes on a shaking table depends strongly on: i) the hardware (mechanical and electrical system), which usually has limitations and therefore can distort the reference signal and; ii) the control system, whose performance is highly sensitive to the adopted control strategy and tuning. The hydraulic and control system supplier, MTS Systems Corporation (MTS) provided several control strategies: i) Three-Variable Control (TVC), also known in control theory literature as state variable control [9], which combines the displacement, velocity and acceleration feedback [1]; ii) Adaptive Harmonic Cancellation (AHC), suitable for sinusoidal signals (fixed frequency or sweep); iii) Adaptive Inverse Control (AIC) [10], suitable for non-sinusoidal signals on linear systems; iv) Amplitude/Phase Control (APC), suitable for sinusoidal signals (fixed frequency and sweep) and predominantly linear systems, and often used in conjunction with AHC to reduce distortions in nonlinear systems; v) Online Iteration (OLI), suitable for nonlinear systems and non-sinusoidal waveforms. This last strategy is the most appropriate for earthquake engineering applications, as structures may exhibit nonlinear behaviour or sustain damage during testing. OLI is currently the control method implemented at LNEC’s ST3D.

Shaking‑table control strategies typically begin with the offline generation of a drive signal using the inverse model of the system and the desired reference motion. This drive signal is then applied to the system, which usually incorporates a linear controller, most commonly a Proportional‑Integral‑Derivative (PID) controller, using actuator feedback signals (displacement, velocity, or acceleration) [1]. This is the basis of the OLI control strategy. Indeed, PID controllers remain widely used in most test systems, including shaking‑table and hybrid testing platforms[11]. Alternative control approaches have been proposed in the literature, including: i) Acceleration Trajectory Tracking Control (ATTC); ii) displacement feedforward-feedback controller using an optimal controller; iii) Model-based feedforward-feedback controller for acceleration tracking [11]; iii) minimal control synthesis (MCS) strategy, which is feedback-feedforward adaptive control strategy including the inverse model in line with the reference signal [1].

Returning to the OLI strategy currently implemented at LNEC’s shaking table, several limitations can be identified: i) the PID controller parameters are fixed and cannot be adapted to the device under test or to the characteristics of the reference signal; ii) the approach requires the offline computation of the drive signal to be fed into the system; iii) the procedure is iterative, often requiring multiple drive‑signal updates to achieve the target response, which increases testing time and may induce additional damage to the specimen. This work presents an alternative control methodology for shaking‑table testing. The system is identified in closed-loop using the two‑stage method, then an optimized controller is synthesized for the identified model and the target earthquake signal. Two linear controllers were considered on this study: i) a PID-F controller, i.e. a PID with a filter on the derivative term, tunned to maximize the bandwidth while maintaining a unitary frequency response function (FRF), using MATLAB’s pidtune function. ii) a LQI controller, i.e. an optimal controller augmented with an integral of the tracking error, whose weighting terms of matrix Q are tunned with the Nelder-Mead simplex optimization algorithm, such that the optimal state feedback law minimizes the mean squared error between the target and output response spectra. The best controller obtained from this process is then selected for implementation. Eight earthquake records were considered to evaluate the performance of the proposed methodology on a small‑scale shaking table equipped with both a PID-F and an optimal LQI controller. Numerical simulations based on the identified physical model demonstrate that these two strategies outperform the OLI control approach currently used at LNEC, with the LQI controller exhibiting the best performance. These results partially validate effectiveness of the proposed methodology and confirm the previous results obtained with the numerical model of the ST1D shaking table [12]. The physical implementation of the controller will be carried out in the near future to fully confirm and validate the approach on the experimental system.

# LNEC’s Shaking Table Test process & Control system

The current control strategy implemented in the LNEC shaking table tests is based on the OLI algorithm. It involves the iterative synthesis of a drive signal with progressively increasing severity, aiming to ensure that the platform accurately reproduces the response spectra of the reference signal, i.e., the target earthquake (Fig. 1).

LNEC’s seismic test process

The process depicted in Fig. 1 consists in testing the structure to a series of test steps starting at a low severity earthquake (scaled earthquake) and then repeated with the same earthquake at higher severity (higher amplitude), and so on until a defined amplitude is reached, which usually corresponds to the original amplitude of the earthquake – reference severity. At each test step the structural damage es accessed and, before starting a test step, an identification of the system's transfer function is carried out, using a limited-band pink noise signal (slope of -10dB/decade). This function is then used in the command signal generation algorithm to generate the drive signal for each test step. Throughout this process, the amplitude of the reference signal increases progressively, and it is expected that the structure will enter a non-linear regime, either due to the behaviour of some of its components or the damage accumulated during the various tests. At each test step, the system's performance is evaluated by comparing the response spectrum of the reference signal with that of the signal measured on the platform (Fig. 2). The response spectrum is a fundamental concept in earthquake engineering and structural dynamics, as it represents the maximum response (such as displacement, velocity, or acceleration) for each natural frequency of a single degree of freedom mass-spring-damper system subjected to the ground motion. It represents how structures respond to a specific earthquake ground-motion record and is used to characterize earthquake severity.

Increasing Severity

Acceleration response spectrum evolution. Four iteratively generated drive signals (for scaled L’Aquila seismic signal) were needed for acceptable performance

Consequently, different earthquake time histories can correspond to the same response spectrum, meaning they have the same severity. However, it is not always possible to achieve satisfactory performance with the command signal synthesized initially, and it may be necessary to generate more than one command signal and carry out additional test steps. This procedure, although necessary to fit the spectrum, it induces additional damage to the structure. Additionally, physical limitations inherent to the equipment, namely with regard to the finite field of displacement, velocity and acceleration, must be duly considered during the tests.

To overcome these limitations a control methodology for shaking-table using an optimized linear controller using the model of the identified system and earthquake recorded is being proposed. The performance of the system will be evaluated with two candidates: a PID-F and an optimal LQI controller.

An experimental facility comprising a small-scale shaking table with dedicated hardware and software was used to validate the methodology. The measurement and control chain is depicted on Fig. 2 and, consists of: i) Shaking Table, including a small-scale shaking table installed on a fixed base, along which it rolls over, driven by a Goodmans model 390A electrodynamic shaker, and a Linear Variable Differential Transformer (LVDT) measures the relative displacement of the platform; ii) Conditioning and Control unit, with a hardware rack comprising all required signal conditioning and control components, including sensor signal conditioners and a National Instruments PXI Real-Time Controller integrated with a CompactRIO system featuring a Virtex-5 FPGA, enabling autonomous real-time execution of control and acquisition tasks; iii) and, Data Acquisition System comprising a host PC, where the experimental setup is configured using the NI LabVIEW environment, including the implementation and supervision of the control strategy, the data acquisition set-up, as well as the measurement and visualization of the experimental results.

Experimental facility

One key advantage of this system framework is the modularity, since both data acquisition system and the condition and control rack can be deployed with other experimental set-ups. In particular, the system can be reused with the ST1D shaking table, where the proposed control strategies will be subsequently validated.

The system under study is a closed-loop control system, schematically illustrated in Fig. 3, comprising: i) a controller, for which both a PID-F controller and an optimal LQI controller are considered in this study; and ii) a controlled system consisting of the shaking-table hardware, including its mechanical and electrical components, together with the physical model. The control objective is that the output signal y, corresponding to the shaking-table motion, tracks the reference signal r as closely as possible.

Shaking table control loop

The control strategy currently implemented at LNEC is the OLI control approach. For the majority of applications, the PID controller parameters are tuned in advance for a specific test configuration and subsequently kept fixed throughout most experimental campaigns.

# Methodology

## Foreword

This section details the methods with which the identification of the plant, tuning of controllers, and evaluation of controller performance was carried out. Firstly, the two-stage identification method is applied in order to address the problem of consistent identification in closed-loop system. The following chapter describes the open-loop driver signal design process currently used at LNEC (which will serve as a baseline for controller performance evaluation), and two proposed developments for the shake table control system. These developments consist on a systematic approach to designing optimized controllers of two different structures: a PID-F controller (for simplicity of implementation on the existing control system software and hardware), and an optimal controller (for greater performance).

## System identification

In order to tune the controllers to the system dynamics, it is necessary to first identify the controlled system. To this end an identification experiment was carried out using a limited-band pink noise driver signal, with the controller C set with just proportional gain. The imperative of running experiments under closed-loop control poses additional challenges for identifying the plant, so, in order to address the particularities of closed-loop identification, the two-stage method was adopted [14].

Firstly, to identify the sensitivity function S0, a linear FIR model with 100 parameters was fitted using MATLAB’s polyest function (the number of parameters was chosen arbitrarily for this first stage).  Then a Hammerstein-Wiener model was defined, using the previously identified FIR model as the linear block, and with the output nonlinearity equal to that of the physical system (saturation of the control action at known values), resulting in a more refined model for the sensitivity function. This function was then used to compute the noise-free control action ur, from which (by being uncorrelated with the output y)  the plant G0 can now be estimated without bias introduced by the controller in the feedback loop.

Block diagram illustrating the two-stage method, with S0 being an  Hammerstein-Wiener model to better represent the real system dynamics with control action saturation

## Control Strategies

LNEC’s driver signal iterative adaptation process: This control technique is known as Online Iteration (OLI) [MTS Systems], as it iteratively modifies the command input on an individual sample-by-sample basis until the system response is close to the target.

This method compensates the next drive signal with a drive correction signal which is computed from the filtering the previous iteration error with the inverse model of the closed-loop frequency response function. The closed-loop frequency response function (FRF) is estimated via a nonparametric identification method by spectral analysis. To this end, the user interface of the driver iteration program allows many different settings for the identification. For the simulations of the OLI method, the identification used the same experimentally acquired data with the band-limited pink noise, and the FRF identification settings had the window type was set to Hanning, number of windows was set to 8192 with 97% overlap, and the FRF mode set to H3. During the driver adaptation process, settings such as correction factor and maximum frequency compensation were adjusted as needed, while the coherence threshold was left at 90%.

PID-F control strategy: Taking into consideration the test procedure at LNEC (Fig. 1), where the identification is carried out each time the target signal is scaled up, or even additional times for various other reasons, one possible straightforward improvement would be to update the controller tune using this information. Another simple enhancement would be to add a filtered derivative term to the PI controller (making it a PID-F), since this should be of simple implementation in the existing version of the control software since it's just an extension of the existing controller structure.

LQI control strategy:	 For the tracking loop shown in Fig. 4, where the state space model of the plant (1) is:

x=A⋅x+B⋅uy= C⋅x+D⋅uA=1.3420.4357-1.3480.5587100001000010B=0.125000C=0.037470.05120.009955-0.003781D=-0.000846	

The optimal state feedback control law u = K [x xi] (where K is 5 element vector, one for each state x, and an additional state xi which is the integral of the tracking error computed via forward Euler formula)  which minimizes the cost function Jut=0∞x⊺Qx+u⊺Ru dtwas computed using MATLAB lqi function, with R=1, and a diagonal Q matrix optimized via Nelder-Mead simplex algorithm [15] through MATLAB’s fminsearch to minimize the acceleration response spectrum MSE computed for 100 logarithmically spaced points between 0.1Hz and 30Hz, which resulted in the approximate weights (rounded to the closest order of magnitude):

Q=10150000010-1000001030000010-2000001016

## System performance evaluation: Response Spectrum

- The performance of each control strategy was evaluated by acceleration and displacement response spectrum mean squared error (MSE). The acceleration response spectrum was computed from 0.1Hz to 30Hz for 500 logarithmically spaced points, and the displacement response spectrum was computed for the same frequencies, however, for only the first 344 points, equating to a range from 0.1Hz to 5.04Hz. The 1DOF mass spring damper system for which the response spectrum was computed consisted of a 1kg mass with a damping ratio ξ = 0.05, for which the spring stiffness and damping were computed to yield the desired natural frequency.

# Results

## Foreword

## Input actions considered: scaled earthquakes

Regarding the input ground motions, a set of recorded earthquake ground-motion time histories covering a wide range of seismic characteristics was considered in this study, namely [13]: i) 1940 Imperial Valley earthquake, El Centro Fault Normal (FN) and Fault Parallel (FP) records; ii) 1992 Erzikan earthquake, Erzikan North-South (NS) and East-West (EW) record; iii) 1995 Kobe earthquake, Kobe NS and EW records; iv) 2009 L’Aquila earthquake; v) 1994 Northridge earthquake, Newhall FN and FP records; vi) 1994 Northridge earthquake, Rinaldi FN and FP records; vii) 1994 Northridge earthquake, Sylmar FN and FP records; viii) 1976 Friuli earthquake, Tolmezzo record. To ensure that the imposed inputs remained within the displacement capacity of the shaking table, appropriate scale factors were applied to the ground-motion signals.

## Summary of results

Acceleration response spectrum for a scaled down Tolmezzo seismic signal

Displacement response spectrum for a scaled down Tolmezzo seismic signal

Acceleration response spectrum MSE for various seismic signals

Displacement response spectrum MSE for various seismic signals

## Conclusions & Outlook

The optimal controller results in a more precise replication of the target response spectra across the whole frequency spectrum. The acceleration response spectrum was particularly improved.

The tuned PID-F improved on the OLI method

The results indicate that the tuned controllers (PID-F \& Optimal) offer superior performance in terms of response spectra compared to the OLI method, while also requiring less user intervention doing so, with optimal control providing the best results in every instance. However, both the performance of the controllers and the adapted drivers depends on identifying the system, and the sensitivity of each method to quality of the identification requires further research.

##### Acknowledgment

The authors acknowledge LNEC for its financial support through the research fellowship under the LNEC Research Programme Modelação do Comportamento Estrutural- Mestr, and Fundação para a Ciência e a Tecnologia (FCT) for its financial support via LAETA (project https://doi.org/10.54499/UID/50022/2025).

##### References

S. J. Chapman, M. S. Williams, A. Blakeborough, “Laboratory testing of structures under dynamic loads: an introductory review”. Philos Trans A Math Phys Eng Sci; 359 (1786): 1651–1669, 15 September 2001. https://doi.org/10.1098/rsta.2001.0860

E. C. Carvalho, “Seismic Testing of Structures”, 11th European Conference on Earthquake Engineering, Balkema, Rotterdam, 1998.

P. X. Candeias, A. A. Correia, G. G. Tekeste, “Shake Table Testing Techniques: Current Challenges and New Trends,” In: Chastre, C., Neves, J., Ribeiro, D., Neves, M.G., Faria, P. (eds) Advances on Testing and Experimentation in Civil Engineering. Springer Tracts in Civil Engineering . Springer, Cham. https://doi.org/10.1007/978-3-031-23888-8_8

T. Emílio, R. T. Duarte, F. Carvalhal, C. O. Costa, C. T. Vaz, M. R. Corrêa, The new LNEC shaking table for earthquake resistance testing”, Memória 757, LNEC, Lisboa, 1989.

R. Bairrão, Duque, Simulador Sísmico Uniaxial - Descrição e operação, Relatório LNEC, 94/1998 - C3ES/SES, LNEC, Lisboa, 1998.

P. Castro, Mesa Vibratória Vertical (Projeto), Relatório LNEC, Obra 76/53/1077, LNEC, Lisboa, Junho de 1978.

F. Carvalhal, Simulation of Seismic Disturbances Using Adaptive Control, MSc Thesis, Victoria University of Manchester, Institute of Science and Technology, October of 1974.

F. Carvalhal, Controlo Adaptativo na Simulação de Solicitações Sísmicas, Tese para Especialista LNEC, STG/DPA, Proc. 76/13/5588, Lisboa, Janeiro de 1979.

K. Ogata, Modern Control Engineering. 5th Edition, Pearson, Upper Saddle River, 2010.

B. Widrow, E. Walach, Adaptive Inverse Control: A Signal Processing Approach, Institute of Electrical and Electronic Engineers, 16 July 2007. DOI:10.1002/9780470231616

G. G. Tekeste, Real-Time Hybrid Simulation including a Shaking Table – Development and Application to Soil-Structure Interaction. PhD Thesis. Univ. de Aveiro, 2021.

A. Henrique, M. A. Botto, F. Oliveira, “Development of control strategies for earthquake reproduction on uniaxial shaking table,” 13º Congresso Nacional de Mecânica Experimental – CNME2026, Faculdade de Engenharia da Universidade do Porto, Porto, 25 a 27 de fevereiro de 2026.

S. Narasimhan, S. Nagarajaiah, E. A. Johnson, H. P. Gavin, “Smart base-isolated benchmark building. Part I: problem definition,” Structural Control and Health Monitoring 13:573–588, 2006. DOI: 10.1002/stc.99

P. M. J. Van Den Hof, R. J. P. Schrama, “An indirect method for transfer function estimation from closed loop data,” Automatica, Volume 29, Issue 6, Pages 1523-1527, 1993. DOI: 10.1016/0005-1098(93)90015-L.

J. C. A. Lagarias, Reeds, M. H. Wright, P. E. Wright. “Convergence Properties of the Nelder-Mead Simplex Method in Low Dimensions.” SIAM Journal of Optimization. Vol. 9, Number 1, 1998, pp. 112–147. DOI:10.1137/S1052623496303470