Development of control strategies for earthquake reproduction on uniaxial shaking table

Desenvolvimento de estratégias de controlo para reprodução de sismos em plataforma sísmica uniaxial

Afonso Costa Henrique1,2 | Miguel Ayala Botto3 | Fernando Pires de Oliveira1

1LNEC – Laboratório Nacional de Engenharia Civil, Portugal          2IST – Instituto Superior Técnico, Portugal      3IDMEC, Instituto Superior Técnico, Universidade de Lisboa, Lisboa, Portugal

Keywords: Earthquake Engineering, Control Systems, System Identification, Closed-loop Identification

Palavras-chave: Engenharia Sísmica, Controlo de Sistemas, Identificação de Sistemas, Identificação em Anel Fechado

Abstract

Despite computational models used in earthquake simulations becoming highly sophisticated and reliable, they still face limitations when complex, nonlinear, or poorly understood structural behaviours are considered.  Moreover, numerical simulations, even with advanced methods, require experimental benchmarks to ensure accuracy, and machine learning approaches also rely on experimental data for training and updating models. Experimental testing provides critical validation and calibration for these models. In summary, earthquake experimental testing remains indispensable for model validation, understanding structural complex behaviours, and ensuring structural safety

This paper analyses control strategies for reproducing earthquake ground motion on the uniaxial shaking table located at Laboratório Nacional de Engenharia Civil, through a computational model of the system. The objective of the controller is to ensure the shaking table reproduces the target earthquake ground motion as accurately as possible. The overall performance is evaluated in terms of response spectra for three distinct control strategies: the currently used method consisting on pseudo-open loop offline adaptive control with an iterative process of generating drive signals and running the tests in closed loop with a PI controller; a tuned PIDF controller (proportional, integral, and filtered derivative) ; and an optimal controller.

The results indicate that both the tuned PIDF and the optimal controllers offer superior performance in terms of response spectra compared to the iterative method, whereas the optimal controller performed the best. The performance of the controllers depends on a reliable dynamic model identification of the uniaxial shaking table, and the sensitivity of each method to the quality of the identification requires further investigation.

# Introduction

Figure 1: Iterative drive signal generation method.

The current shaking table earthquake reproduction process at LNEC consists on iteratively synthesizing a drive signal (figure 1), distinct from the target motion, such that the shaking table replicates the response spectra of the target, i.e. the earthquake ground motion.

At the end of each test, the system's performance is evaluated by comparing the response spectrum of the reference signal with that of the output signal measured on the platform. However, it is not always possible to achieve satisfactory performance with the command signal synthesized initially, and it may be necessary to generate more than one drive signal and carry out additional tests at the same amplitude (figure 2). This procedure, although necessary in some situations, should be avoided as it induces additional damage to the structure.

Figure 2: LNEC's test process. New driver signals are generated (loop highlighted in yellow box) till the shaking platform produces the desired response spectrum (plot on the right). Following this, the target severity is increased and the process is repeated.

# Objectives

The objective of this work is to implement a control strategy to ensure that the uniaxial shaking table reproduces the earthquake target signal as accurately as possible, eventually improving the current iterative method.

Control Strategies: Performance evaluation in simulation environment

Two controllers, a PIDF and an optimal controller, were tuned following the identification performed before each test.  Afterwards, for each of the controllers, the system performance in reproducing the target motion on the shaking table was evaluated (figure 3). The effectiveness of the proposed controllers was verified for a set of benchmark earthquakes.

Conclusions

Figure 3: Proposed test process with controller tunning at every severity level, and replacing iterative driver design entirely.

The results indicate that the proposed controllers offer superior performance in terms of response spectra compared to the iterative method, while also requiring less user intervention in the process, with the optimal control strategy providing the best results in every instance. The performance of the tuned controllers was found to outperform the iterative method even after three iterations, achieving satisfactory performance on the first run, which is highly desirable. However, the performance of each method relies on accurate system identification, and the sensitivity of each method to this factor requires further investigation.